공통규격서

Commodity Description

물품번호 Item No.	관세분류번호 HSK. No.	정부물품분류번호(8자리) Korean Government Commodity Classification Code(eight-digit)	품 명 Description	단 위 Unit	수 량 Quantity
1	9027.80.2090	41115413	장시간세포관찰분석시스템 (Long Term Live Cell Screening Analysis System)	Set	1

Ⅰ. 용도(End-user's Use)

1) kinetic proliferation and cell growth monitoring

2) kinetic Cytotoxicity, Apoptosis, ReporterGene, Migration, Invasion assay etc.

3) image analysis of cell morphology during cell incubation

Ⅱ. 장비의 구성(Configurations of Goods)

1. Main Body(주장비)

- Microscope Gantry                                            1 ea

- Controller Unit                                                1 ea

- Base software	1 ea

2. Accessories

- Migration and Invasion assay kit                               1 ea

- Marking Tool and Pen	1 ea

- Laptop PC                                                    1 ea

Ⅲ. 성능 및 규격(Performance and Specification)

1. Microscope Gantry

1) Observation method : Phase contrast and Fluorescence (Green & Red) or more

2) Magnification:

- 4x / 10x / 20x objectives or more

- User can change them in less a minute.

3) Native Image Size : 1344 x 1024 pixels  or more

4) Image Resolution

- 4x : 3.05 um/pixel or more

- 10x : 1.22 um/pixel or more

- 20x : 0.61 um/pixel or more

5) Exported Image Format: JPEG, PNG, TIFF, RAW or more

6) Exported Movie Format: WMV, AVI or more

7) Fluorescence Excitation/Emission :

- Green Channel Excitation 440-480 nm

- Green Channel Emission 504-544 nm

- Red Channel Excitation 565-605 nm

- Red Channel Emission 625-705 nm or more

8) Vessel types : Microplate, Dish, Slide or more

9) Vessel Capacity: up to 6 x 96 (or 384)well microplates at once or 1536well microplate

10) Environmental control : Operating in CO2 incubator or Optional system like

chamber for Temperature, Humidity, CO2 condition should be attached on

main body.

2. Controller Unit

1) Internal Data Storage:  9 Terabyte RAID Hard Drive, 12 Gbyte RAM,

12 Core processor or more

2) Network Connection: Ethernet 10/100/1000 Mbps or more

3) Power: 100-240VAC 50/60Hz, 3.5 Amps max@120V , 1.75Amps max@240V

3. Base Software

1) Image acqusition Interval : minimum 1 min. / maximun 24 hours or more.

2) Movie file : Crop, Zoomed movie available (Time , Scale bar selection)

3) Data : Images, Movie file, Value, Graph.

4) Data Export : Excel, word etc.

5) Compatible with Microsoft Windows XP, Vista 7

6) Application module : Cell Proliferation, Cytotoxicity, etc.

7) Phase metrics and filters :

- Cell or Colony Confluency (%)

- Cell or Colony Count

- Cell or Colony Area (um2)

- Cell or Colony Shape (Eccentricity)

8) Green and Red Fluorescence Image Acquisition

9) Fluorescence Migration And Invasion

10) Analysis : Fluorescence Object Counting

4. Cell Migration and Invasion Assay kit

1) HD Phase contrast and Fluorescence modes

2) includes Woundmaker to creat wound area  on 96well plate in just seconds.

3) Measure both 2D migration and 3D Invasion on the same 96 well plate at 4x,

10x or 20x.

4) Automated analysis algorithms for Relative Wound Density, Wound confluence,

Wound width.

5. Marking Tool and Pen : one marking tool holder with one pen

6. Laptop PC : 17" monitor size, QuadCore, 4Gbyte RAM or more,  500Gbyte HDD or more.

Ⅳ. 기타 조건(Remarks)

1. This installation and operation should be performed by the bidder/supplier's responsibility at end user's

site

2. One year warranty service should be provided by supplier after the date of installations